## Slide 1
Low temperature sterilization of flexible endoscopes
Protection from detrimental impacts of the sterilant during the process
Dr. Tronje von dem Hagen, Lübeck
tronje@vondemhagen.com
An increasing number of flexible endoscope applications ask for sterile conditions.
Practice from the last years in reprocessing using low temperature sterilization processes with chemical sterilizing agents show, that repeated process conditions finally can impair the functional or mechanical integrity of the endoscope due to accumulation of small deteriorating impacts.
Respective limitations of sterilization cycles by the endoscope manufacturers and extra costs for frequent inspection, maintenance and repair ask for technical improvements; some of them have been analyzed, implemented , and successfully tested.

## Slide 2
7-10 oct. 2015
Corrosion and detrimental residues due to ingress of aggressive sterilant into the inside
Rupture of protective flexible tube at distal end due to missing ETO cap in a vacuum process
Typical damages following (frequent) low temperaturer sterilization cycles (1)

## Slide 3
7-10 oct. 2015
Damage of light guide fibres, e.g. due to penetration of liquids through defective front lens sealing
Deterioration of adhesive around distal front lens due to chemical incompatibility to the sterilant
Typical damages following frequent low temperaturer sterilization cycles (2)

## Slide 4
7-10 oct. 2015
(Cobra Vision, Richard Wolf)
Endoscope outer surfaces (need to be reprocessed / sterilized):
Specific design aspects of endoscopes (1)

## Slide 5
7-10 oct. 2015
(Cobra Vision, Richard Wolf)
Endoscope outer surfaces (need to be reprocessed / sterilized):
Control head with handles
Umbilical cable /universal cord
Connectors for media supplies
Ports for insertion of instruments
Insertion tube
Optical components (e.g. front lenses)
Operating channels (inner surfaces)
Specific design aspects of endoscopes (1)

## Slide 6
7-10 oct. 2015
Specific design aspects of endoscopes (2)
Endoscope housing Inside
(hermetically sealed, no need to be reprocessed):

## Slide 7
7-10 oct. 2015
Specific design aspects of endoscopes (2)
Endoscope housing Inside
(hermetically sealed, no need to be reprocessed):
Manipulation mechanics
Angulation control strings
Light guides
Channel tubes (outer surfaces)
Valve housings
Electrical connectors
Electronic parts

## Slide 8
7-10 oct. 2015
Specific aspects to be considered for reprocessing
Possible defects, mistakes, and residual risks
Sealing of endoscope housing defective (leak): reprocessing chemicals may penetrate and can damage inner parts (residues, chemical reaction)
Vacuum phases during sterilization generate a high pressure difference between inside and outside; internal overpressure (> 300mbar) may burst the very thin and flexible protective tube at the distal end of the insertion tube.
Missing ETO-Cap during sterilization using processes with vacuum phases (user mistake); rupture of protective tube
Sneaky buildup of defects and residuals due to individually neglectable impacts at each sterilization cycle, becoming detectable and relevant for operation only upon accumulation after 100 cycles or more.
Leak test is mandatory before reprocessing, test device to be connected to leak test port (internal overpressure max. 300 mbar)
The „ETO-Cap“ to be connected to the leak test port opens the spring loaded valve and allows pressure equilibration between inside and outside of the endoscope housing
Training of the staff; in case of rupture: Endoscope out of order for repair at customer service
Material compatibility tests in co-operation between the manufacturers of the endoscope and the sterilizer; Limitation of number of reprocessing cycles until scheduled maintenance.
100 cycles

## Slide 9
7-10 oct. 2015
Example 2
Note: An endoscope might be used and reprocessed some hundreds of times a year.
Example 3
Consequences for routine operation (current situation)
Example 1
100 cycles,
EtO cap
inspect

## Slide 10
7-10 oct. 2015
Critical positions for ingress of sterilizing agents are:
The distal front of the insertion tube endoscope manufacturers shall improve compatibility of adhesives and sealing coatings (permanent task)
a) integrate a non-return valve (or any other) into the ETO- Cap to allow evacuation, but to block ingress of aggressive chemical processing (sterilizing) agents
(to be designed by the sterilizer manufacturer)
b) integrate a non-return valve (or any other) into the leak test port
(to be designed by the endoscope manufacturer)
The Leak test port intentionally opened by the EtO-cap during vacuum sterilization cycles
Technical options to improve:

## Slide 11
7-10 oct. 2015
Design sketch of a Sterilization Valve combined with the EtO-Cap (schematic)
EtO-CAP
Leak test port
Symbol
Non-return valve on top
New Sterilization Valve

## Slide 12
7-10 oct. 2015
Application of a sterilization valve 1. Example: STERIVIT- high/low pressure process
3
time
0
1
2
4
5
desorption phase
ca. 8 - 10 hrs
pressure
ca. 10 - 12 hrs
[bar]
Holding time
2,7 bar
6,5 bar
6
Use of a sterilization valve not recommended for EtO
(keep the EtO-Cap)
very high pressure differences
no restrictions are known for EtO, no concerning information from practice
NO

## Slide 13
7-10 oct. 2015
Application of a sterilization valve 2. Example: STERRAD 100 NX process
Use of a non return valve recommended for H2O2 (with or without plasma)
no overpressure during process
relevant detrimental impact of H2O2 experienced in practice
ventilation after aeration phase?
Process pressure inside the sterilizer chamber (outside endoscope housing)
Pressure inside endoscope housing (valve adjusted to open/close at ca. 50 mbar)
Pressure [mbar]
aeration
H2O2 diffusion
H2O2 diffusion
plasma
plasma
0.1
1
10
100
1000
aeration
conditionning
50 mbar
time
- 960 mbar
Air injection
YES

## Slide 14
7-10 oct. 2015
Application of a sterilization valve 3. Example: LTSF-process (EN 14180)
Use of a sterilization valve recommended for LTSF
no overpressure during process
relevant detrimental impact of formaldehyde /water in practice
ventilation after aeration phase?
Process pressure inside the sterilizer chamber (outside endoscope housing)
Pressure inside endoscope housing (valve adjusted to open/close at ca. 170 mbar)
time
Pressure
Po
1013
200
50
conditionning
Desorption & drying
Holding time
170 mbar
- 790 mbar
formaldehyde + steam steam
air
YES

## Slide 15
7-10 oct. 2015
Ventilation after aeration phase (end of cycle)?
Double non-return valve
(single) non-return valve
Retaining vacuum inside the endoscope housing after end of cycle and storage at these conditions for some time may create problems for specific designs of the endoscopes (e.g. mechanical push buttons under a flexible sealing plastic cap).
Options:
Use of a double non-return valve
Manual lifting of the non-return valve
Use of a servo valve (venting the inner volume at ambient pressure)
Solution is influenced by the process parameters, to be investigated by the sterilizer manufacturer.

## Slide 16
7-10 oct. 2015
Measured cycle profile for LTSF process 60°C matachana 130 LF / WEBECO FA95 using double non return valve LTSF-valve (+170 / -170 mbar)
Chamber temperature
Chamber pressure
Pressure inside endoscope housing
Residual vacuum: 170 mbar

## Slide 17
7-10 oct. 2015
Type O
Type W
Type S
Type F
Type P
Example from practice (specific LTSF-valve [double non-return], matachana-WEBECO)
Endoscope specific Coonector
Sterilization Valve

## Slide 18
7-10 oct. 2015
Low temperature sterilization of flexible endoscopes
Protection from detrimental impacts of the sterilant during the process
Summary
Critical design characteristics of flexible endoscopes have been analyzed and technical options could be found to eliminate identified impacts.
Technical solutions for most relevant LTS processes as applied in the market and possible improvements in general have been discussed.
Some design options and practical solutions have been presented and their functionality for LTSF sterilization has been demonstrated by experimental data using a double non-return valve available in the market.
Experience from recent years has proven proper function and reliability of this protective valve, which is recommended by an increasing number of endoscope manufacturers.
Use of the valve could verify significant reduction of maintenance and repair costs.

## Slide 19
7-10 oct. 2015
Low temperature sterilization of flexible endoscopes
Protection from detrimental impacts of the sterilant during the process
Dr. Tronje von dem Hagen, Lübeck
tronje@vondemhagen.com
Thank you very much
for your attention