## Slide 1
ENDOSCOPE STORAGE/DRYING CABINET: IMPORTANCE OF QUALIFICATIONS
Lionel PINEAU PhD, BIOTECH-GERMANDE, FRANCE

## Slide 2: ENDOSCOPE REPROCESSING PROCEDURE
Patient
Cleaning/disinfection
Drying
Storage
Reprocessing
<12h

## Slide 3: THE RISK ASSOCIATED WITH STORAGE
"Accurate endoscope drying is crucial, whereas a humid environment facilitates microbial growth during storage. The final drying steps greatly reduce the risk of remaining pathogens as well as the possibility of recontamination of the endoscope by waterborne microorganisms such as Pseudomonas spp."
Kovaleva J. Transmission of Infection by Flexible Gastrointestinal Endoscopy and Bronchoscopy Clinical Microbiology Reviews 2013

## Slide 4: THE RISK ASSOCIATED WITH STORAGE
Lawrence F. Muscarella, Am J Gastroenterol 2006;101:2147–2154)

## Slide 5: ENDOSCOPE STORAGE CABINET
DSC800 Soluscope
AS300 HYSIS
EDC Olympus
EDC Plus Olympus
ED200 Steelco
ED250 Steelco
Scope Store PT LTE
DSC800 Soluscope
AS300 HYSIS
EDC Olympus
EDC Plus Olympus
ED200 Steelco
ED250 Steelco
Scope Store PT LTE

## Slide 6: THE EUROPEAN STANDARD EN 16442
  France has proposed a new work item based on the existing French standard,
  The new work item has been accepted on May 2011,
  Working Group 8 of TC102 was in charge to draw up the draft European standard,
  The document was submitted to CEN members for formal vote in September 2014,
  The European standard EN 16442 was published in March 2015.

## Slide 7
ESC are designed to provide a controlled environment for the storage of endoscope(s) (with or without channels) and, if necessary, drying of the endoscope(s).
The controlled environment provided by the ESC, ensures that during storage, there is no deterioration of the microbiological quality of the endoscope. The drying function is intended to supplement, if necessary, any drying conducted during automated or manual processing of the endoscope.
THE EUROPEAN STANDARD EN 16442

## Slide 8: EN 16442: TESTING FOR CONFORMITY
Performance qualification:
process of obtaining and documenting evidence that the equipment, as installed and operated in accordance with operational procedures, consistently performs in accordance with predetermined criteria and thereby yields product meeting its specification.
Type tests
(prototype)
Routine Testing
Factory
On site
Tests performed by or on behalf of the manufacturer
Tests performed by or on behalf of the user
Opertionnal Qualification
Performance Qualification
Work tests
(Serial production)
Installation Qualification

## Slide 9: EN 16442: TESTING FOR CONFORMITY
Performance qualification tests shall be performed on at least one model of each relevant endoscope type tests group as defined in Annex F and the choice shall be made according to the endoscopes available on site.
or
Alternatively performance tests may be performed with one model of each endoscope product family (see Annex G) for which the same ESC connector set is used, if it is demonstrated that the endoscopes selected are the most challenging.

## Slide 10: EN 16442: ENDOSCOPE TYPE TEST GROUP
It is impossible to repeat all type tests on every flexible endoscope which can be processed in the ESC.
The objective is to group the devices based on the similarities of design and perform the tests only on a representative sample of each of these endoscope type test groups.

## Slide 11: EN 16442: ENDOSCOPE TYPE TEST GROUP
General as well as specific features of flexible endoscopes shall be considered when clustering endoscopes in endoscope type test groups i.e. design, ports, connectors, channel separators, port closures, internal non-return valves, restrictions in channels, internal connection between channels.
Channel separators

## Slide 12: FRENCH GUIDELINES (1, 2)
For performance qualification, a significant number of endoscopes (one model of each endoscopes family and a number of endoscopes corresponding to at least one third of the total capacity of the ESC) is sampled after storage.
Endoscope family: endoscopes with the same internal design (number of channels, channel diameter,…) and same ESC connectors (all characteristics which may have an influence on the ESC efficacy are identical).
(1) Recommandations de bonnes pratiques d'utilisation des enceintes de stockage d'endoscopes thermosensibles (ESET). SF2H, SFED. 2011, 7 pages.
(2) Haut Conseil de la santé publique. AVIS. Enceintes de stockage d’endoscopes thermosensibles (ESET) 26 juin 2013

## Slide 13: PQ(a): ENDOSCOPE TO BE TESTED
Endoscopes available on site | Endoscope Type Test Group
A | TTG1
B | TTG1
C (b) | TTG1
D | TTG2
E | TTG3
F | TTG3
G | TTG4
H | TTG3
… | 
ESC connector | Endoscope product family | Endoscopes available on site
C1 | Colonoscope | A
C1 | Colonoscope | B
C1 | Colonoscope | C
C2 | Colonoscope | D
C3 | Gastroscope | E
C3 | Gastroscope | F
C4 (c) | Gastroscope | G
C3 | Gastroscope | H
 | … | …
Selected endoscopes for PQ
C
D
G
…
(a): Performance Qualification. (b): longest endoscope. (c): C4 = C3 + 1 extra connection
Endoscope Type Test Group (Annex F)
Endoscope product family (Annex G)

## Slide 14: EN 16442: TEST PROGRAM
a Alternative test procedure for the parametric operational and/of performance qualification of storage cabinet is given in Annex D of this standard
b These tests are optional if the parametric operational and performance qualification procedures are followed (see Annex D)
X – Recommended,
B – Not recommended,
O– Optional

## Slide 15: EN 16442: SURFACE CONTAMINATION
6.5.1.1 The efficacy of the recommended cleaning/disinfection procedure shall be verified by determining the contamination level with contact agar from 4 zones in the chamber located as follows:
— 2 zones that could be in contact with the endoscope during storage,
— 1 zone at other location in the chamber of the ESC,
— 1 zone at the bottom.
Method: The zones tested shall have a surface area of around 25 cm2, with Trypticase soya agar to determine the presence of bacteria and filamentous fungi by incubation at (30±1)°C for at least five days.
http://www.biomerieux-usa.com/

## Slide 16: EN 16442: SURFACE CONTAMINATION
6.5.1.2 Testing shall be performed at the end of the recommended time before application of the recommended cleaning/disinfection procedure.
6.5.2 Acceptance criteria
The contamination levels identified shall be less than 25 CFU/25 cm2.
NOTE A contamination level lower than 25 CFU is not considered to be satisfactory if the microorganisms recovered are considered to be pathogenic for the intended use of the device. This situation can require further investigation to identify the type and source of contamination.

## Slide 17: EN 16442: ENDOSCOPES
Verification of the ability of the ESC to maintain the quality of the endoscopes
E.2.2.1 The endoscopes intended to be stored in the storage cabinet are cleaned and disinfected, stored in the storage cabinet according to the instructions for use for the maximum storage time specified for the type of endoscope being tested and sampled.
E.2.2.2 Samples taken from the distal ends of the endoscopes are split into two equal volumes. Each volume is analysed by membrane filtration and incubated on Tryptone soya agar and Sabouraud dextrose agar with chloramphenicol and incubated for 5 days at (30 ± 1) °C.

## Slide 18: EN 16442: ENDOSCOPES
E.2.2.3 Following incubation, the colonies formed are counted and the results are expressed in number of viable microorganisms per endoscope (X) taking into account the recovery ratio of the sampling method.
E.2.3 Acceptance criteria
The results are deemed to be acceptable if contamination of the internal channels of the endoscopes is less than 25 CFU/endoscope.
NOTE 2 A contamination level lower than 25 CFU/endoscope is not considered to be satisfactory if the microorganisms recovered are considered to be pathogenic for the intended use of the device. This situation can require further investigation to identify the type and source of contamination.

## Slide 19: ENDOSCOPE SAMPLING SOLUTION
Different sampling solutions are described in the literature for endoscope sampling:
Sterile distilled water,
Saline solution(1) [NaCl (0.9%)],
FHM(1,2) [Peptone (0.01%), Tween 80 (0.1%), NaCl (0.043%), Phosphate (0.095%)],
DNP(3) [Peptone (0.01%), Tween 80 (0.3%), Lecithin (0.01%), Histidin chlorhydrate (0.01%) NaCl (0.043%), KH2PO4 (0.036%), Na2HPO4, 2H20 (0.072%)],
T+Thio(4) [Tween 80 (0.3%), Lecithin (0.03%), L-Histidin (0.01%) Sodium thiosulfate (0.05%)].
Eléments d’assurance qualité en hygiène relatifs au contrôle microbiologique des endoscopes et à la traçabilité en endoscopie. Conseil supérieur d’hygiène publique de France. March 2007.
NF EN ISO 11737-1
DNP: Neutralizing Pharmacopoeia Diluent.
Sampling solution recommended by Biotech-Germande

## Slide 20: ENDOSCOPE SAMPLING SOLUTION
Chart: Sampling solutions
| DNP | FHM | NaCl 0,9% | T+Thio
1st sampling | 65.5 | 50.32 | 0 | 76.5
2nd sampling | 21.58 | 26.4 | 1.98 | 15.32
Validation of endoscope sampling solutions (1)
The recovery ratio of the sampling method was determined by repeated rounds of sample collection according to ISO 11737-1 annex C1.(2, 3)
76,5%
(1) RICHARD M, LUU DUC D, PINEAU L. Efficacy of recovery solutions for endoscopes sampling : a comparative study. SHEA 19th Annual Scientific Meeting, San Diego, March 21st 2009
(2) ISO = International Organization for Standardization.
(3) ISO 11737-1 annex C1 Guideline. Available at: http://www.iso.org/iso/home/store/catalogue_tc/catalogue_detail.htm?csnumber=46116. Accessed 10/15/12.
< 2%

## Slide 21: INITIAL ENDOSCOPE QUALITY
EN 16442 E.2.2.1 NOTE 1 All endoscopes available in the unit where the storage cabinet is installed and that are intended to be stored in the ESC cabinet should be sampled before installation of the ESC (initial assessment) in order to be able of determining the source of the contamination in the event that a routine test is positive.
French guidelines: The installation of the ESC shall be preceded by:
An audit of the endoscope reprocessing procedure,
A control of the microbiological quality of all endoscopes, intended to be stored in the ESC (initial assessment) within the month preceding the installation.

## Slide 22: EN 16442: DRYING TESTS
6.4.3 Exterior surface drying
Load the storage cabinet, following the instructions for use. Carry out a normal operating cycle. Immediately following the recommended drying period, use a piece of anhydrous copper (II) sulphate paper (or crepe paper) to test the presence of moisture on the load exterior surfaces (copper (II) sulphate paper from white or light blue to deep blue) as evidence of residual water).
6.4.4 Channel drying
At the end of the drying phase, the endoscope shall be removed from the cabinet. Direct the distal end of the endoscope towards a horizontal piece of anhydrous copper (II) sulphate paper at a distance of 50 mm to 100 mm and flush medical grade air at a positive pressure.

## Slide 23: EN 16442: AIR QUALITY
6.6.1 Air cleanliness
If a specific level of particulate cleanliness for the air used in the storage cabinet is claimed/required, then measurements of the specific level of particle cleanliness shall be carried out to check whether the particulate cleanliness class achieved matches the manufacturer's claims (without endoscopes, doors closed, in the centre of the storage cabinet chamber after a stabilization period of not less than 15 min and not more than 20 min ).
EN ISO 14644-1, Cleanrooms and associated controlled environments — Part 1: Classification of air cleanliness (ISO 14644-1)

## Slide 24: EN 16442: AIR QUALITY
5.2.2.1 Airborne microbial contamination
Air inside the ESC and flowing through the channels of the endoscope shall be of a microbiological quality which will not impair the quality of the load. Tests shall be done according to Annex C .
Active sampling: < 100 CFU/m3 (Class 8 of ISO 14644-1).
  1 m3 of air is sampled on Trypticase soya agar.
  The agar is then incubated at (30 ± 1) °C for 5 days.
Sedimentation: < 50 CFU/m3 (Class 8 of ISO 14644-1)
4 Trypticase Soya agar plates are placed on the floor of each
chamber of the storage cabinet and left for 1 hour. The agar
cultures are then incubated at (30 ± 1) °C for 5 days

## Slide 25: EN 16442: THERMOMETRIC TEST
6.9.1 Procedure:
Fit the temperature sensors as follows:
— one sensor on the control head of the endoscope or,
— one sensor in at least one endoscope channel at the distal end or,
—on the outer surface of the endoscope.
Record the temperature during a drying phase or a limited time frame during the storage cycle.
6.9.2 Acceptance criteria
The results are considered to be satisfactory when the temperatures recorded are within the specified temperature band for the corresponding stage.

## Slide 26: EN 16442: CHANNEL AERATION TEST
Procedure
Two endoscopes, from the list of compatible endoscopes and that is used by the user, one that has the highest air demand and one that has the lowest air demand are connected to the ESC.
The flow of air in each tubing of the ESC connector is blocked one by one and the moment the storage cabinet indicates a fault is recorded.
Acceptance criteria
Verify that the ESC indicate a fault and that the moment the fault is indicated corresponds to the setting of the channel aeration control

## Slide 27: ROUTINE TESTING
Endoscope: An endoscope representative of each endoscope family stored in the ESC shall be sampled quarterly and each endoscope at least one time a year,
Air cleanliness: required If a specific level of particulate cleanliness for the air used in the storage cabinet is claimed,
Airborne microbial contamination : Sedimentation, 1 agar plate/1 hour, < 25 CFU/m3 (at least quarterly),
Surface contamination: 5 sampling points, < 25 CFU/25 cm2 (at least quarterly).

## Slide 28: CONCLUSIONS
The use of en endoscope storage cabinet ensures that during storage there is no deterioration of the microbiological quality of the endoscope.
Confirmation that the ESC is consistent with the essential requirements of EN 6442 is essential but not sufficient.
Considering that it si not possible to repeat type tests for every flexible endoscopes intended to be processed in the storage cabinet, performance qualification and routine sampling are critical to ensure that the equipment, as installed and operated in accordance with operational procedures, consistently performs in accordance with predetermined criteria on all endoscopes intended to be processed in the unit where the ESC is installed.

## Slide 29: CONCLUSIONS
THANK YOU VERY MUCH…